## Analog Circuits Reference
[7 empty] | Constant Voltage | [5 empty] | Constant Current
[7 empty] | Resolution | V Range | Zener Max |  |  | Divider R |  |  | Ref Res | Constant I
[7 empty] | 15 | 5.01 | 5.02 |  |  | 75000 |  |  | 330 | 0.02
Res | 10 | 12 | 14 | 15 | 16 |  | Level % | Step # | V_Reading | R Result | mA | mW | Level % | Step # | V_Reading | R Result | mW | R mA
Steps | 1024 | 4096 | 16384 | 32768 | 65536 |  | 0.001 | 0 | 0 | 0 | 0.0668 | 0 | 0.01 | 3 | 0.00045868 | 0.022936 | 0.0091736 | 0.019999
V | 0.00025 | 6.25e-05 | 1.5625e-05 | 7.8125e-06 | 3.9063e-06 |  | 0.005 | 2 | 0.00030579 | 4.5688 | 0.066796 | 2.0425e-05 | 0.1 | 33 | 0.0050455 | 0.25247 | 0.10091 | 0.019985
mV | 0.25 mV | 0.063 mV | 0.016 mV | 0.008 mV | 0.004 mV |  | 0.01 | 3 | 0.00045868 | 6.8534 | 0.066794 | 3.0637e-05 | 0.5 | 164 | 0.025074 | 1.2585 | 0.50149 | 0.019924
[7 empty] | 0.05 | 16 | 0.0024463 | 36.566 | 0.066767 | 0.00016333 | 1 | 328 | 0.050149 | 2.5266 | 1.003 | 0.019848
[7 empty] | 0.1 | 33 | 0.0050455 | 75.4564 | 0.066733 | 0.0003367 | 2 | 655 | 0.10014 | 5.0844 | 2.0029 | 0.019697
@5 mA supply means a resistor of 1000 ohms, which gives 5 mV | [6 empty] | 1 | 328 | 0.050149 | 756.7972 | 0.066133 | 0.0033165 | 5 | 1638 | 0.25044 | 13.0158 | 5.0088 | 0.019241
[7 empty] | 2 | 655 | 0.10014 | 1526.645 | 0.065467 | 0.0065562 | 10 | 3277 | 0.50103 | 27.1095 | 10.0206 | 0.018482
[7 empty] | 5 | 1638 | 0.25044 | 3938.08 | 0.063467 | 0.015895 | 25 | 8192 | 1.2525 | 77.2931 | 25.05 | 0.016205
|  | Divider Resistor | 120 |  |  |  | 10 | 3277 | 0.50103 | 8315.4564 | 0.060133 | 0.030128 | 50 | 16384 | 2.505 | 201.8681 | 50.1 | 0.012409
|  | Vref | 0.256 |  |  |  | 25 | 8192 | 1.2525 | 24933.643 | 0.050133 | 0.062792 | 75 | 24576 | 3.7575 | 436.2269 | 75.15 | 0.0086136
[7 empty] | 50 | 16384 | 2.505 | 74701.7893 | 0.033467 | 0.083834 | 90 | 29491 | 4.509 | 711.5917 | 90.1794 | 0.0063365
[7 empty] | 75 | 24576 | 3.7575 | 223217.8218 | 0.0168 | 0.063125 | 95 | 31130 | 4.7596 | 853.4134 | 95.1912 | 0.0055771
[7 empty] | 90 | 29491 | 4.509 | 661746.5192 | 0.0068002 | 0.030662 | 98 | 32113 | 4.9099 | 958.6469 | 98.1971 | 0.0051217
[7 empty] | 95 | 31130 | 4.7596 | 1370636.895 | 0.0034656 | 0.016495 | 99 | 32440 | 4.9599 | 997.9282 | 99.197 | 0.0049701
[7 empty] | 98 | 32113 | 4.9099 | 3343222.7234 | 0.0014657 | 0.0071962 | 99.9 | 32735 | 5.005 | 1035.4783 | 100.0991 | 0.0048335
[7 empty] | 99 | 32440 | 4.9599 | 6184463.4087 | 0.00080039 | 0.0039698
[7 empty] | 99.5 | 32604 | 4.9849 | 10659305.5024 | 0.00046673 | 0.0023266 | 78 | 25559 | 3.9078 | 479.002 | 78.1559 | 0.0081582
[7 empty] | 99.9 | 32735 | 5.005 | 24949141.4981 | 0.00020021 | 0.001002
[7 empty] | 99.99 | 32765 | 5.0095 | 35923809.4891 | 0.00013917 | 0.00069718
Voltage Divider
|  | Current Limiter |  | Sense
|  | 500000 |  | 44000
| Vin | R1 | Vdrop | R2 | Vdrop | R3 | Gnd
| 50 | 500000 | 26.0536 | 44000 | 23.9464 | 500000 | 0
[6 empty] | V across Sense |  |  | Approx V Scale Factor
[6 empty] | 2.1073 |  |  | 22.7273

## Calibrations
DMM Calibration Template
Reference and Setup |  |  |  |  | Subject Meter
[5 empty] | EEPROM 1-1 |  |  | EEPROM 1-2 |  |  | EEPROM 1-3 |  |  | EEPROM 1-4 |  |  | EEPROM 1-5 |  |  | EEPROM 1-5 |  |  | Giga
LETTER | PCB Row | Nominal Value | Actual | Difference | Reading | Correction Factor | Pasteable Text | Reading | Correction Factor | Pasteable Text | Reading | Correction Factor | Pasteable Text | Reading | Correction Factor | Pasteable Text | Reading | Correction Factor | Pasteable Text | Reading | Correction Factor | Pasteable Text | Reading | Correction Factor | Pasteable Text
A | 1 | 0.5 | 0.491 | 0.982 | 0.5 | 0.982 | CF_A = 0.982; | 0.47492 | 1.0339 | CF_A = 1.0339; | 0.4976 | 0.9867 | CF_A = 0.9867; | 0.4996 | 0.9828 | CF_A = 0.9828; | 0.50314 | 0.9759 | CF_A = 0.9759; | 0.50314 | 0.9759 | CF_A = 0.9759; | 0.5111 | 0.9607 | CF_A = 0.9607;
B | 2 | 1 | 0.997 | 0.997 | 1.0173 | 0.98 | CF_B = 0.98; | 0.98665 | 1.0105 | CF_B = 1.0105; | 0.99628 | 1.0007 | CF_B = 1.0007; | 1.0012 | 0.9958 | CF_B = 0.9958; | 1.0007 | 0.9963 | CF_B = 0.9963; | 1.0007 | 0.9963 | CF_B = 0.9963; | 1.0124 | 0.9848 | CF_B = 0.9848;
C | 3 | 4.7 | 4.651 | 0.98957 | 4.7352 | 0.9822 | CF_C = 0.9822; | 4.6706 | 0.9958 | CF_C = 0.9958; | 4.657 | 0.9987 | CF_C = 0.9987; | 4.6516 | 0.9999 | CF_C = 0.9999; | 4.6584 | 0.9984 | CF_C = 0.9984; | 4.6584 | 0.9984 | CF_C = 0.9984; | 4.6652 | 0.997 | CF_C = 0.997;
D | 4 | 10 | 10.021 | 1.0021 | 10.272 | 0.9756 | CF_D = 0.9756; | 10.1489 | 0.9874 | CF_D = 0.9874; | 10.05 | 0.9971 | CF_D = 0.9971; | 10.0425 | 0.9979 | CF_D = 0.9979; | 10.0485 | 0.9973 | CF_D = 0.9973; | 10.0485 | 0.9973 | CF_D = 0.9973; | 10.063 | 0.9958 | CF_D = 0.9958;
E | 5 | 33 | 32.91 | 0.99727 | 30.749 | 1.0703 | CF_E = 1.0703; | 34.0393 | 0.9668 | CF_E = 0.9668; | 33.07 | 0.9952 | CF_E = 0.9952; | 33.0626 | 0.9954 | CF_E = 0.9954; | 33.0648 | 0.9953 | CF_E = 0.9953; | 33.0648 | 0.9953 | CF_E = 0.9953; | 33.128 | 0.9934 | CF_E = 0.9934;
F | 6 | 100 | 99.49 | 0.9949 | 98.221 | 1.0129 | CF_F = 1.0129; | 98.2002 | 1.0131 | CF_F = 1.0131; | 99.88 | 0.9961 | CF_F = 0.9961; | 99.8439 | 0.9965 | CF_F = 0.9965; | 99.8372 | 0.9965 | CF_F = 0.9965; | 99.8372 | 0.9965 | CF_F = 0.9965; | 99.84 | 0.9965 | CF_F = 0.9965;
G | 7 | 330 | 331.5 | 1.0045 | 329.4 | 1.0064 | CF_G = 1.0064; | 330.0741 | 1.0043 | CF_G = 1.0043; | 332.5 | 0.997 | CF_G = 0.997; | 332.18 | 0.998 | CF_G = 0.998; | 332.38 | 0.9974 | CF_G = 0.9974; | 332.38 | 0.9974 | CF_G = 0.9974; | 332.09 | 0.9982 | CF_G = 0.9982;
H | 8 | 1000 | 1000.1 | 1.0001 | 1000.492 | 0.9996 | CF_H = 0.9996; | 999.7791 | 1.0003 | CF_H = 1.0003; | 995.6776 | 1.0044 | CF_H = 1.0044; | 996.7263 | 1.0034 | CF_H = 1.0034; | 996.7503 | 1.0034 | CF_H = 1.0034; | 996.7503 | 1.0034 | CF_H = 1.0034; | 997.3469 | 1.0028 | CF_H = 1.0028;
I | 9 | 3000 | 3002 | 1.0007 | 3004.04 | 0.9993 | CF_I = 0.9993; | 3004.8323 | 0.9991 | CF_I = 0.9991; | 2992.8354 | 1.0031 | CF_I = 1.0031; | 2995.6272 | 1.0021 | CF_I = 1.0021; | 2996.7166 | 1.0018 | CF_I = 1.0018; | 2996.7166 | 1.0018 | CF_I = 1.0018; | 2998.3438 | 1.0012 | CF_I = 1.0012;
J | 10 | 10000 | 9928 | 0.9928 | 9945.82 | 0.9982 | CF_J = 0.9982; | 9944.4669 | 0.9983 | CF_J = 0.9983; | 9905.3516 | 1.0023 | CF_J = 1.0023; | 9916.1719 | 1.0012 | CF_J = 1.0012; | 9905.1982 | 1.0023 | CF_J = 1.0023; | 9905.1982 | 1.0023 | CF_J = 1.0023; | 9923.2988 | 1.0005 | CF_J = 1.0005;
K | 11 | 33000 | 32970 | 0.99909 | 33081.32 | 0.9966 | CF_K = 0.9966; | 33002.2086 | 0.999 | CF_K = 0.999; | 32825.832 | 1.0044 | CF_K = 1.0044; | 32868.0273 | 1.0031 | CF_K = 1.0031; | 32883.457 | 1.0026 | CF_K = 1.0026; | 32883.457 | 1.0026 | CF_K = 1.0026; | 32913.1484 | 1.0017 | CF_K = 1.0017;
L | 12 | 100000 | 100090 | 1.0009 | 100973.83 | 0.9912 | CF_L = 0.9912; | 100404.0401 | 0.9969 | CF_L = 0.9969; | 99177.039 | 1.0092 | CF_L = 1.0092; | 99363.992 | 1.0073 | CF_L = 1.0073; | 99464.57 | 1.0063 | CF_L = 1.0063; | 99464.57 | 1.0063 | CF_L = 1.0063; | 99742.68 | 1.0035 | CF_L = 1.0035;
M | 13 | 333000 | 326400 | 0.98018 | 336925.69 | 0.9688 | CF_M = 0.9688; | 327829.5946 | 0.9956 | CF_M = 0.9956; | 319441.656 | 1.0218 | CF_M = 1.0218; | 320741.844 | 1.0176 | CF_M = 1.0176; | 321344 | 1.0157 | CF_M = 1.0157; | 321344 | 1.0157 | CF_M = 1.0157; | 323954.875 | 1.0075 | CF_M = 1.0075;
N | 14 | 1000000 | 1038500 | 1.0385 | 1144097.75 | 0.9077 | CF_N = 0.9077; | 1078753.7799 | 0.9627 | CF_N = 0.9627; | 968296.4375 | 1.0725 | CF_N = 1.0725; | 978735.125 | 1.0611 | CF_N = 1.0611; | 980444.875 | 1.0592 | CF_N = 1.0592; | 980444.875 | 1.0592 | CF_N = 1.0592; | 1014431.75 | 1.0237 | CF_N = 1.0237;
O | 15 | 4700000 | 4705000 | 1.0011 | 6443228.5 | 0.7302 | CF_O = 0.7302; | 6787841.5249 | 0.6932 | CF_O = 0.6932; | 3834936.25 | 1.2269 | CF_O = 1.2269; | 4227492 | 1.113 | CF_O = 1.113; | 4321954 | 1.0886 | CF_O = 1.0886; | 4321954 | 1.0886 | CF_O = 1.0886; | 4190996.25 | 1.1226 | CF_O = 1.1226;
[5 empty] | Voltage |  |  | Voltage |  |  | Voltage |  |  | Voltage |  |  | Voltage |  |  | Voltage |  |  | Voltage
[5 empty] | Reading | True Value | Constant | Reading | True Value | Constant | Reading | True Value | Constant | Reading | True Value | Constant | Reading | True Value | Constant | Reading | True Value | Constant | Reading | True Value | Constant
[7 empty] | #DIV/0! |  |  | #DIV/0! |  |  | #DIV/0! |  |  | 46.4676 | 9.3672 | 9.37 | 46.4815 | 10.087 | 9.364 | 46.4162 |  |  | #DIV/0!
0
0

## BOM
Components
Reference | Value | Original | Datasheet | Footprint | Qty | DNP | Description | Function | Purchase Link | Cost | Qty2 | Cost each | Cost / DMM | QTY On Hand | Builds Possible
C1 | 0.47uF | N | ~ | Capacitor_SMD:C_0805_2012Metric_Pad1.18x1.45mm_HandSolder | 1 |  | Unpolarized capacitor | 9v supply filter1 | https://www.digikey.com/en/products/detail/murata-electronics/GRM21BR71H474KA88L/702592 | 7.7 | 100 | 0.077 | 0.077 | 10 | 10
C16,C19 | 22uF | N | ~ | Capacitor_SMD:C_0805_2012Metric_Pad1.18x1.45mm_HandSolder | 2 |  | Unpolarized capacitor | Boost Output Smoothing | https://www.digikey.com/en/products/detail/samsung-electro-mechanics/CL21A226MAYNNNE/10479857 | 2.49 | 30 | 0.083 | 0.166 | 19 | 9
C18 | 1uF | N | ~ | Capacitor_SMD:C_0805_2012Metric_Pad1.18x1.45mm_HandSolder | 1 |  | Unpolarized capacitor | R Circuit Smoothing | https://www.digikey.com/en/products/detail/samsung-electro-mechanics/CL21B105KBFNNNE/3886687 | 1.6 | 100 | 0.016 | 0.016 | 100 | 100
C3,C14 | 2.2uF | N | ~ | Capacitor_SMD:C_0805_2012Metric_Pad1.18x1.45mm_HandSolder | 2 |  | Unpolarized capacitor | -- mixed values -- | https://www.digikey.com/en/products/detail/samsung-electro-mechanics/CL21B225KOFNNNE/3886820 | 1.86 | 30 | 0.062 | 0.124 | 20 | 10
C4,C17 | 0.1uF | N | ~ | Capacitor_SMD:C_0805_2012Metric_Pad1.18x1.45mm_HandSolder | 2 |  | Unpolarized capacitor | -- mixed values -- | https://www.digikey.com/en/products/detail/samsung-electro-mechanics/CL21B104KACNNNC/3886757 | 8 | 100 | 0.08 | 0.16 | 0 | 0
C9,C10,C15 | 4.7uF | N | ~ | Capacitor_SMD:C_0805_2012Metric_Pad1.18x1.45mm_HandSolder | 3 |  | Unpolarized capacitor | -- mixed values -- | https://www.digikey.com/en/products/detail/samsung-electro-mechanics/CL21A475KAQNNNG/3894439 | 8 | 100 | 0.08 | 0.24 | 15 | 5
D1 | LED | Y | ~ | LED_SMD:LED_Luminus_MP-3030-1100_3.0x3.0mm | 1 |  | Light emitting diode | LED SMD Option | https://www.digikey.com/en/products/detail/luminus-devices-inc/MP-3030-110H-40-95/9593557 | 2.86 | 20 | 0.143 | 0.143 | 15 | 15
D4,D5 | D_Schottky | N | ~ | Diode_SMD:D_SOD-123 | 2 |  | Schottky diode | Overvoltage Proection | https://www.digikey.com/en/products/detail/onsemi/BAT54T1G/918319 | 4.08 | 20 | 0.204 | 0.408 | 90 | 45
F1 | Fuse | N | ~ | Fuse:Fuseholder_Clip-5x20mm_Keystone_3512P_Inline_P23.62x7.27mm_D1.02x2.41x1.02x1.57mm_Horizontal | 2 |  | Fuse | Dead Short Protection | https://www.digikey.com/en/products/detail/bel-fuse-inc/FC-211-BRIGHT-TIN/622587 | 9.49 | 100 | 0.0949 | 0.1898 |  | 0
J11 | I_Input | N | ~ | Connector_PinHeader_2.54mm:PinHeader_1x03_P2.54mm_Horizontal | 1 |  | Generic connector, single row, 01x03, script generated (kicad-library-utils/schlib/autogen/connector/) |  | https://www.amazon.com/Chanzon-Straight-Connector-Breadboard-Electronic/dp/B09MYMD4YS?crid=108K7I4JQX3JW&dib=eyJ2IjoiMSJ9.HwO1Ueh1Z7svwElyoT6FBehn2EReHR_6JHvtM8pZx4bGLaJcsmiEA7Yv0aOdh1CI9u3rDtrjwU2sqMTEwzScEulyWwc3395JvuqxKdl_kjELh35buIqS7VWH920V8xZgmvY7h7gPR1R9a9J_6yqRTumY7kyiL1ajMPT4hnPR_TQYG4zaH-sxvrl1N2IQJSBLK1bzKIcLpz4UK2dlLv7F3nDID7Jq6p_9LmaMXV5X9Cw.FhF8Cc_MXqxzGcmAwWrgMNoDQAozI9rdi6q7z3Lwgcs&dib_tag=se&keywords=1x4%2Bpin%2Bheader&qid=1749847050&sprefix=1x4%2Bpin%2B%2Caps%2C113&sr=8-8&th=1 | 6.99 | 50 | 0.1398 | 0.1398 |  | 0
J12 | Conn_01x02_Pin | N | ~ | Connector_PinHeader_2.54mm:PinHeader_1x02_P2.54mm_Horizontal | 1 |  | Generic connector, single row, 01x02, script generated |  | https://www.amazon.com/Chanzon-Straight-Connector-Breadboard-Electronic/dp/B09MYL5M4K?crid=108K7I4JQX3JW&dib=eyJ2IjoiMSJ9.HwO1Ueh1Z7svwElyoT6FBehn2EReHR_6JHvtM8pZx4bGLaJcsmiEA7Yv0aOdh1CI9u3rDtrjwU2sqMTEwzScEulyWwc3395JvuqxKdl_kjELh35buIqS7VWH920V8xZgmvY7h7gPR1R9a9J_6yqRTumY7kyiL1ajMPT4hnPR_TQYG4zaH-sxvrl1N2IQJSBLK1bzKIcLpz4UK2dlLv7F3nDID7Jq6p_9LmaMXV5X9Cw.FhF8Cc_MXqxzGcmAwWrgMNoDQAozI9rdi6q7z3Lwgcs&dib_tag=se&keywords=1x4%2Bpin%2Bheader&qid=1749847050&sprefix=1x4%2Bpin%2B%2Caps%2C113&sr=8-8&th=1 | 6.99 | 50 | 0.1398 | 0.1398 |  | 0
J14,J15 | Barrel_Jack | N | ~ | Connector_BarrelJack:BarrelJack_Horizontal | 2 |  | DC Barrel Jack |  | https://www.amazon.com/California-JOS-Breadboard-Friendly-Mounting-Female/dp/B09XH1XJZT?crid=1K1D82IK4Q41W&dib=eyJ2IjoiMSJ9._2Ev6i_66TQNaOQKz-SGN1wTjIQ6wwy57JnL7cNVBemvDJgUfzpUD3fx6PFTxcIK-GoqXpWRN4pfPPl28PJR1N6vit5kTOYn54SZZqF7cklC1C3kX1nAdIa0m207sNCWut0M8Vvqe45ziJ4pUTsQAJPDHLMRla1M5KOXfYHf2hkDHePylkLYbZkIlgwszZVMLkYYhjuQrnyRCkWF0CJBIQ3falYAdmGJePqv1A8qFXQ.1ciiigbhrIeOXutM04Xnra9lNSGn8i4Z1vPy07TjTZ4&dib_tag=se&keywords=dc%2Bbarrel%2Bjack&qid=1749856721&sprefix=dc%2Bbarrel%2Bjac%2Caps%2C194&sr=8-5&th=1 | 8.99 | 10 | 0.899 | 1.798 |  | 0
J4 | IR Connection | N | ~ | Library:PinHeader_1x03_P2.54mm_Vertical_NoBox | 1 |  | Generic connector, single row, 01x03, script generated (kicad-library-utils/schlib/autogen/connector/) |  | https://www.amazon.com/Chanzon-Straight-Connector-Breadboard-Electronic/dp/B09MYMD4YS?crid=108K7I4JQX3JW&dib=eyJ2IjoiMSJ9.HwO1Ueh1Z7svwElyoT6FBehn2EReHR_6JHvtM8pZx4bGLaJcsmiEA7Yv0aOdh1CI9u3rDtrjwU2sqMTEwzScEulyWwc3395JvuqxKdl_kjELh35buIqS7VWH920V8xZgmvY7h7gPR1R9a9J_6yqRTumY7kyiL1ajMPT4hnPR_TQYG4zaH-sxvrl1N2IQJSBLK1bzKIcLpz4UK2dlLv7F3nDID7Jq6p_9LmaMXV5X9Cw.FhF8Cc_MXqxzGcmAwWrgMNoDQAozI9rdi6q7z3Lwgcs&dib_tag=se&keywords=1x4%2Bpin%2Bheader&qid=1749847050&sprefix=1x4%2Bpin%2B%2Caps%2C113&sr=8-8&th=1 | 6.99 | 50 | 0.1398 | 0.1398 |  | 0
J9 | Out to Screen | N | ~ | Connector_PinSocket_2.54mm:PinSocket_1x04_P2.54mm_Vertical | 1 |  | Generic connector, single row, 01x04, script generated |  | https://www.amazon.com/Chanzon-Straight-Connector-Breadboard-Electronic/dp/B09MYG647K?crid=108K7I4JQX3JW&dib=eyJ2IjoiMSJ9.HwO1Ueh1Z7svwElyoT6FBehn2EReHR_6JHvtM8pZx4bGLaJcsmiEA7Yv0aOdh1CI9u3rDtrjwU2sqMTEwzScEulyWwc3395JvuqxKdl_kjELh35buIqS7VWH920V8xZgmvY7h7gPR1R9a9J_6yqRTumY7kyiL1ajMPT4hnPR_TQYG4zaH-sxvrl1N2IQJSBLK1bzKIcLpz4UK2dlLv7F3nDID7Jq6p_9LmaMXV5X9Cw.FhF8Cc_MXqxzGcmAwWrgMNoDQAozI9rdi6q7z3Lwgcs&dib_tag=se&keywords=1x4%2Bpin%2Bheader&qid=1749847050&sprefix=1x4%2Bpin%2B%2Caps%2C113&sr=8-8&th=1 | 6.99 | 50 | 0.1398 | 0.1398 |  | 0
L1 | 2.2uH | N | ~ | Inductor_SMD:L_Murata_DFE201610P | 1 |  | Inductor |  | https://www.digikey.com/en/products/detail/murata-electronics/DFE201612E-2R2M-P2/9815901 | 4.08 | 20 | 0.204 | 0.204 | 13 | 13
Q1,Q3,Q4,Q5 | MMBT3904 | Y | https://www.onsemi.com/pdf/datasheet/pzt3904-d.pdf | Package_TO_SOT_SMD:SOT-23 | 4 |  | 0.2A Ic, 40V Vce, Small Signal NPN Transistor, SOT-23 | -- mixed values -- | https://www.digikey.com/en/products/detail/diotec-semiconductor/MMBT3904/13163698 | 3.98 | 100 | 0.0398 | 0.1592 | 20 | 5
Q2 | SI2302 MOSFET | Y | http://www.diodes.com/assets/Datasheets/DMG2302U.pdf | Package_TO_SOT_SMD:SOT-23 | 1 |  | 4.2A Id, 20V Vds, N-Channel MOSFET, SOT-23 | R Circuit Range Switch | https://www.digikey.com/en/products/detail/mcc-micro-commercial-components/SI2302-TP/1793243 | 4.14 | 20 | 0.207 | 0.207 | 16 | 16
R1 | 500 | N | ~ | Resistor_SMD:R_0805_2012Metric_Pad1.20x1.40mm_HandSolder | 1 |  | Resistor, US symbol | I Limit LED | https://www.digikey.com/en/products/detail/yageo/RC0805FR-07499RL/727986 | 1.12 | 100 | 0.0112 | 0.0112 | 10 | 10
R12 | 100k | N | ~ | Resistor_SMD:R_0805_2012Metric_Pad1.20x1.40mm_HandSolder | 1 |  | Resistor, US symbol | MOSFET Bleeder1 | https://www.digikey.com/en/products/detail/yageo/RC0805JR-07100KL/728239 | 0.82 | 100 | 0.0082 | 0.0082 |  | 0
R17,R18 | 499.0K | Y | ~ | Resistor_SMD:R_0805_2012Metric_Pad1.20x1.40mm_HandSolder | 2 |  | Resistor, US symbol | -- mixed values -- | https://www.digikey.com/en/products/detail/panasonic-electronic-components/ERA-6AED4993V/19246650 | 1.86 | 20 | 0.093 | 0.186 | 4 | 2
R2,R13,R15,R16,R25,R26 | 1K | N | ~ | Resistor_SMD:R_0805_2012Metric_Pad1.20x1.40mm_HandSolder | 6 |  | Resistor, US symbol | -- mixed values -- | https://www.digikey.com/en/products/detail/susumu/RR1220P-102-D/432291 | 2.68 | 100 | 0.0268 | 0.1608 | 50 | 8
R5 | 62 | Y | ~ | Resistor_SMD:R_0805_2012Metric_Pad1.20x1.40mm_HandSolder | 1 |  | Resistor, US symbol | R Circuit I Set | https://www.digikey.com/en/products/detail/panasonic-electronic-components/ERA-6AEB620V/1465719 | 6.93 | 50 | 0.1386 | 0.1386 | 40 | 40
R6 | 330 | Y | ~ | Resistor_SMD:R_0805_2012Metric_Pad1.20x1.40mm_HandSolder | 1 |  | Resistor, US symbol | 330 R Circuit I Bleeder | https://www.digikey.com/en/products/detail/stackpole-electronics-inc/RNCF0805DTE330R/6241750 | 1.95 | 50 | 0.039 | 0.039 | 30 | 30
R7, R8, R27 | 10K | N | ~ | Resistor_SMD:R_0805_2012Metric_Pad1.20x1.40mm_HandSolder | 3 |  | Resistor, US symbol | Current Limiting | https://www.digikey.com/en/products/detail/yageo/RC0402JR-0710KL/726418 | 0.61 | 100 | 0.0061 | 0.0183 |  | 0
R9,R14 | 22.0K | Y | ~ | Resistor_SMD:R_0805_2012Metric_Pad1.20x1.40mm_HandSolder | 2 |  | Resistor, US symbol | -- mixed values -- | https://www.digikey.com/en/products/detail/susumu/RR1220P-223-D/432323 | 3.94 | 100 | 0.0394 | 0.0788 | 70 | 35
SW1,SW2 |  | N |  | Button_Switch_THT:SW_PUSH_6mm | 2 |  | Push button switch, generic, two pins |  | https://www.amazon.com/QTEATAK-Momentary-Tactile-Button-Switch/dp/B07VSNN9S2?dib=eyJ2IjoiMSJ9.FvCpFhIAwff79PSVuLNg4c-H-9hpghcEQ-6QLY25pwepWKd3n5-D71iGk6yl2ME_fzV1E8SOmR5-U5uyA2CSgsVP3s6UGNiSQTqADLaHJMP9DPg9Xv7onAJ7Gfm0CLAEFArvfK5LPHqkz9KAPNITbvSyIl1qRHeBpsE37EfgfKUpVneUuXfB_X4D4kuIDiCblIoa_mk2OuFtueni4HTiT6t7W2UJDl5aSpkyQHXc5cA.M_SF_56GBtkHBHo6oYCqsEENudhJ13kk5NEz1TyXsFc&dib_tag=se&keywords=6mm+button&mfadid=adm&qid=1749855961&sr=8-9 | 6.99 | 150 | 0.0466 | 0.0932 |  | 0
SW3 | SW_SPDT | N | ~ | Button_Switch_THT:SW_Slide_SPDT_Angled_CK_OS102011MA1Q | 1 |  | Switch, single pole double throw | On / Off | https://www.digikey.com/en/products/detail/c-k/OS102011MA1QN1/1981430 | 0.54 | 10 | 0.054 | 0.054 | 6 | 6
SW4 | SW_Push | N | ~ | SMD | 3 |  | Push button switch, generic, two pins |  | https://www.digikey.com/en/products/detail/c-k/PTS636SM43SMTR-LFS/10071723 | 4.06 | 25 | 0.1624 | 0.4872 | 20 | 6
U1 | LM317_TO-220 | Y | http://www.ti.com/lit/ds/symlink/lm317.pdf | Package_TO_SOT_SMD:SOT-223-3_TabPin2 | 1 |  | 1.5A 35V Adjustable Linear Regulator, TO-220 |  | https://www.digikey.com/en/products/detail/texas-instruments/LM317MDCYR/454187 | 9.13 | 25 | 0.3652 | 0.3652 | 5 | 5
U10 | LM4040DBZ-2.5 | Y | http://www.ti.com/lit/ds/symlink/lm4040-n.pdf | Package_TO_SOT_SMD:SOT-23 | 1 |  | 2.500V Precision Micropower Shunt Voltage Reference, SOT-23 | V Reference Control | https://www.digikey.com/en/products/detail/texas-instruments/LM4040CIM3X-2-5-NOPB/367111 | 11.99 | 25 | 0.4796 | 0.4796 | 20 | 20
U11 | LM4040DBZ-5 | Y | http://www.ti.com/lit/ds/symlink/lm4040-n.pdf | Package_TO_SOT_SMD:SOT-23 | 1 |  | 5.000V Precision Micropower Shunt Voltage Reference, SOT-23 | R Circuit Reference | https://www.digikey.com/en/products/detail/texas-instruments/LM4040CIM3-5-0-NOPB/305136 | 11.99 | 25 | 0.4796 | 0.4796 | 20 | 20
U3 | ADS1115IDGS | Y | http://www.ti.com/lit/ds/symlink/ads1113.pdf | Package_SO:TSSOP-10_3x3mm_P0.5mm | 1 |  | Ultra-Small, Low-Power, I2C-Compatible, 860-SPS, 16-Bit ADCs With Internal Reference, Oscillator, and Programmable Comparator, VSSOP-10 |  | https://www.digikey.com/en/products/detail/texas-instruments/ADS1115IDGSR/2231567 | 33.31 | 10 | 3.331 | 3.331 | 1 | 1
U7 | MAX1683 | Y | https://www.analog.com/media/en/technical-documentation/data-sheets/MAX1682-MAX1683.pdf | Package_TO_SOT_SMD:SOT-23-5_HandSoldering | 1 |  | IC REG CHARGE PUMP 2VIN SOT23-5 | Voltage Doubler | https://www.digikey.com/en/products/detail/microchip-technology/MCP1253-33X50I-MS/529833 | 18.5 | 10 | 1.85 | 1.85 | 2 | 2
U9 | TPS613222ADBV | Y | http://www.ti.com/lit/ds/symlink/tps61322.pdf | Package_TO_SOT_SMD:SOT-23-5 | 1 |  | 1.8A Step-Up Converter, 5V Output Voltage, 0.9-5.5V Input Voltage, SOT-23-5 |  | https://www.digikey.com/en/products/detail/texas-instruments/TPS613222ADBVR/8638376 | 6.77 | 25 | 0.2708 | 0.2708 | 20 | 20
[11 empty] | Component Cost: |  | 12.5027
Modules
U6 | XIAO RA4M1 |  |  |  | 1 |  |  |  |  | 12.59 | 3 | 4.1967 | 4.1967
Battery |  |  |  |  | 1 |  |  |  | https://www.amazon.com/1100mAh-Rechargable-Protection-Insulated-Development/dp/B0867KDMY7?ref_=pd_ci_mcx_mh_pe_im_d1_hxwPPE_sspa_dk_det_cav_p_0_0&pd_rd_i=B0867KDMY7&pd_rd_w=K8iwI&content-id=amzn1.sym.57b80066-10e8-4be7-a5f2-ce3f1faa4959&pf_rd_p=57b80066-10e8-4be7-a5f2-ce3f1faa4959&pf_rd_r=DPGH2XTEAH7AXE98T8RT&pd_rd_wg=R2tXj&pd_rd_r=bcfd4e9a-22c6-4970-ba41-9f2861f04913&th=1 | 18.99 | 4 | 4.7475 | 4.7475
Screen |  |  |  |  | 1 |  |  |  | https://www.amazon.com/dp/B0CN373JF4?ref=ppx_yo2ov_dt_b_fed_asin_title | 12.88 | 5 | 2.576 | 2.576
[11 empty] | Module Costs |  | 11.5202
[8 empty] | PCBA Cost: |  | Price | Qty | Avg Cost
[9 empty] | Assembly | 92 | 50 | 1.84
[9 empty] | Shipping | 76.59 | 50 | 1.5318
[9 empty] | PCB | 110 | 50 | 2.2
[9 empty] | Stencil | 12 | 50 | 0.24
[9 empty] | Modules | 576.0083 | 50 | 11.5202
[9 empty] | Components | 625.135 | 50 | 12.5027
[9 empty] | Total | 1491.7333 | 50 | 29.8347
Elecrow Quote
[10 empty] | 650 | 50 | 13
[9 empty] | Modules | 576.0083 | 50 | 11.5202
[9 empty] | Total | 1226.0083 | 50 | 24.5202